# Question 1

Load Voltage

The time-dependence of the load voltage produced by the Marx generator is shown in figure 1. The load voltage has a peak voltage of 493kV, a rise time of 98.4ns and an estimated bandwidth of 3.56MHz. To measure this impulse using a voltage sensor with an accuracy of 2% requires the voltage sensor to have a rise time which is five times faster than that of the load voltage. This means that the sensor is required to have a rise time of 19.7ns and a bandwidth of 17.8MHz.

Figure 1: Marx Generator Discharge Load Voltage

Voltage Divider

The low voltage arm of the divider will carry the voltage to be measured by the oscilloscope; therefore to reduce the effect of mismatched loads distorting the signal a simple design for the low voltage arm has been chosen. The characteristics of this design are shown in Table 1; the value chosen for Zc is 50Ω.

| Z1 | Z2 | Coax Cable | Z | Tlv |
| --- | --- | --- | --- | --- |
| Zc | Zc | yes | Zc/2 | 0 |

Table 1: Low Voltage Arm Characteristics

The oscilloscope that will measure the Marx generator output is only capable of measuring a maximum of 80V, therefore, using the attenuation factor equation and the value of the peak Marx generator output the high voltage arm resistance is required to be in the region of 308kΩ. The high voltage arm is surrounded by polythene with a surface breakdown voltage of 5kV/cm requiring the high voltage column to be 98.5cm tall to avoid the voltage breaking across the polythene.

Using the Ayrton-Perry method to reduce the inductance of the system requires the resistance of each wire to double, so that the parallel combination of the resistances is equal to the required value. For a mandrel radius of 5cm and using wire with a diameter of 24µm will require 404m of wire at 643 turns over the height of the resistive column. The following code calculates the change in temperature of the chosen wire using the total energy in the Marx generator pulse. The time over which the energy is calculated was chosen as the time at which the load voltage reduced to zero. This results in a temperature change of approximately 80K. Increasing the diameter of the wire would increase the amount of wire needed but would also decrease the temperature change.

%Energy

tmax = 130*10^-6;

Q = ((tmax*(max(VLoad)^2))/(2*Zhv));

%Wire Mass

wire_vol = (pi*(wire_radius^2))*twirelength; %m^3

wire_mass = wire_vol*evan_dens; %kg

%Temperature

deltaT = Q./(wire_mass.*evan_heat)

To minimise the stray capacitance the grading structure shown in figure 2 is used. This involves two circular rings, made of aluminium, on both the high voltage and low voltage ends of the voltage divider. This method creates an equipotential field along the height of the resistive column with a capacitance, calculated using Maxwell SV, of 16pF.

Figure 2: Field Controlling Structure for Voltage Divider.

Figure 3 shows that a metallic object can be placed up to 80cm away from the field without causing much disturbance to the electric field. Therefore the high voltage connecting lead has a length of 80cm and is made from wire with a diameter of 10mm. To reduce the inductance of the connecting wire 4 are connected in parallel with a total inductance of 200nH. The total inductance for the system is calculated using the equation for a rectangular arrangement and produced a value of 3.3µH.

Figure 3: Field Controlling Structure with Foreign Metallic Object

Using the calculated inductance and capacitance of the system the transfer function of the system is a general second order system in the form;

Fp= 15.3e-17p2+8.7e-9+1

Therefore the response time of the circuit can be expressed as below for an ideal damping ratio of 10% overshoot;

Tconn= rdc where;

rd= 1.2lc

The following code calculates the rise time of the system as a whole using the ideal damping ratio of 10% to produce the total rise time of the system as 10.9ns. Comparing this value to the required rise time of 19.7ns shows that the system is capable of measuring the voltage from the Marx generator system.

rd_sys = 1.2*sqrt(indSys./cHV);

rd_connector = 1.2*sqrt(indConnector./cHV);

Ttotal = sqrt(((rd_sys.*cHV)+(rd_connector.*cHV))^2)

As mentioned above the system transfer function is that of a second order system, therefore the overshoot of the system can be expressed as below;

overshoot= e-π*ζ1-ζ2

For an ideal overshoot of 10% the damping factor ζ can be expressed as;

e-π*ζ1-ζ2=0.1

-π*ζ1-ζ2=log0.1

-π*ζ=log0.1*1-ζ2

π2*ζ2=log0.12*(1-ζ2)2

ζ2(π2+log0.12)=log0.12

ζ2=log0.12(π2+log0.12)

ζ=log0.1(π2+log0.12)

An alternative method of designing the low voltage arm would be to choose the design shown in table 2. From this design the low voltage arm would have a negative rise time; this would reduce the overall rise time of the system.

| Z1 | Z2 | Coax Cable | Z | Tlv |
| --- | --- | --- | --- | --- |
| R+l | Zc | yes | (R+lp)ZcR+lp+Zc | -lZcR(R+Zc) |

Table 2: Alternative Low Voltage Arm Characteristics

# Question 2

Figure 4 and 5 show the current discharge in the RLC circuit ad it’s differential.

Figure 4: RLC Current Discharge

Figure 5: RLC Differential Current Discharge

Coil Design

The Rogowski coil was chosen to have characteristics of; major radius 2cm, minor radius 2mm, made with a wire of radius 1mm with a total of 10 turns. The cable impedance was chosen as 50Ω. The following code calculates the self-inductance and resistance for the described characteristics of the Rogowski coil and produces results of 38nH and 6.2mΩ. Figure 6 shows the time rate of change of the magnetic flux in the Rogowski coil.

%Inductance

Induct_pre = u0.*minor_radius.*N;

Induct_sum0 = 0.0007*((log((2*pi*major_radius)/p))^0);

Induct_sum1 = 0.1773*((log((2*pi*major_radius)/p))^1);

Induct_sum2 = -0.0322*((log((2*pi*major_radius)/p))^2);

Induct_sum3 = 0.00197*((log((2*pi*major_radius)/p))^3);

Induct_sum = Induct_sum0 + Induct_sum1 + Induct_sum2 + Induct_sum3;

Induct_pre2 = (((pi*minor_radius)/p)+(log((2*p)/(wire_diameter)))-(5/4)-Induct_sum);

RCoil_inductance = Induct_pre*Induct_pre2

%Resistance

freq = 1/(2*pi*sqrt(Lb*Cb));

RCoil1 = N/(pi*wire_diameter);

RCoil2 = sqrt((copper_rho*pi*freq*u0)*((p^2)+((2*pi*minor_radius)^2)));

RCoil_resistance = RCoil1*RCoil2

Figure 6: Differential Magnetic Flux

The following coil calculates the current and differential current in the Rogowski coil. Figures 7 and 8 show the current and differential current in the coil.

Rrt = RCoil_resistance + Rcvr;

Lr = RCoil_inductance;

omega = sqrt((1./(Lb.*Cb))-((Rb./(2.*Lb))^2));

rogowski_current = @(x) exp((Rrt./Lr).*x).*(k.*(Vo./(omega.*Lb).*((omega.*exp(-1.*(Rb./(2.*Lb)).*x).*cos(omega.*x))-((Rb./(2.*Lb)).*exp(-1.*(Rb./(2.*Lb)).*x).*sin(omega.*x)))));

t2 = 0:1*10^-10:10*10^-9;

RIntegral = zeros(0,length(t2));

for n=1:length(t2)

RIntegral(n) = integral(@(x)rogowski_current(x),0,t2(n));

end

Rogowski_Current = (exp(-(Rrt./Lr).*t2)./Lr).*RIntegral;

dRogowski_Current_prefix = (exp(-(Rrt./Lr).*t2)./Lr).*exp((Rrt./Lr).*t2).*(k.*(Vo./(omega.*Lb).*((omega.*exp(-1.*(Rb./(2.*Lb)).*t2).*cos(omega.*t2))-((Rb./(2.*Lb)).*exp(-1.*(Rb./(2.*Lb)).*t2).*sin(omega.*t2)))));

dRogowski_Current = dRogowski_Current_prefix + (RIntegral.*((-Rrt.*exp((-Rrt./Lr).*t2))./(Lr^2)));

Figure 7: Rogowski Current

Figure 8: Differential Rogowski Current

Using the same code as above and the total resistance and inductance of the coil, figure 9 demonstrates that the designed coil is acting as a differential coil for 0.5ns. Figure 10 shows the output voltage of the Rogowski coil.

Figure 9: Sensor Characteristics

Figure 10: Rogowski Coil Output Voltage

Figure 11 shows the output voltage after the integrator circuit with a resistance of 50mΩ and a capacitance of 500nF.

Figure 10: Integrator Output Voltage

# Question 3

Figures 11 and 12 show the current discharge for the over-damped RLC circuit ad it’s differential.

Figure 11: RLC Current Discharge

Figure 12: Differential RLC Current Discharge

The Rogowski coil was chosen to have characteristics of; major radius 10cm, minor radius 2.5mm, made with a wire of radius 1.5mm with a total of 15 turns. The cable impedance was chosen as 5Ω. The following code calculates the self-inductance and resistance for the described characteristics of the Rogowski coil and produces results of 93.5nH and 8.2mΩ. Figure 13 shows the time rate of change of the magnetic flux in the Rogowski coil. Figures 14 and 15 show the current and differential current in the Rogowski coil.

%Inductance

Induct_pre = u0.*minor_radius.*N;

Induct_sum0 = 0.0007*((log((2*pi*major_radius)/p))^0);

Induct_sum1 = 0.1773*((log((2*pi*major_radius)/p))^1);

Induct_sum2 = -0.0322*((log((2*pi*major_radius)/p))^2);

Induct_sum3 = 0.00197*((log((2*pi*major_radius)/p))^3);

Induct_sum = Induct_sum0 + Induct_sum1 + Induct_sum2 + Induct_sum3;

Induct_pre2 = (((pi*minor_radius)/p)+(log((2*p)/(wire_diameter)))-(5/4)-Induct_sum);

RCoil_inductance = Induct_pre*Induct_pre2;

%Resistance

freq = 1/(Rb.*Cb);

% freq = 1./(Rb.*Cb);

RCoil1 = N/(pi*wire_diameter);

RCoil2 = sqrt((copper_rho*pi*freq*u0)*((p^2)+((2*pi*minor_radius)^2)));

RCoil_resistance = RCoil1*RCoil2;

Figure 13: Differential Magnetic Flux

Figure 14:

# Appendix A

Q1.m

%% Fast Tansient Sensors - Q1 - Coursework 3

% B126949 - Tom Young

%% Pre - Cursor

clear all

clc

%% Constants

%Global constants

u0 = 4*pi*10^-7;  % H/m

e0 = 8.85*10^-12; % F/m

%Wire constants

evan_rho = 1.38*10^-6; %Ohm/m

evan_heat = 435; %J/kg/K

evan_dens = 8100; %kg/m^3;

%Generator Constants

Ns = 20;

CStage = 40*10^-9;

V0 = 25*10^3;

RStage = 10*10^-3;

RLoad = 10*10^3;

CLoad = 250*10^-9;

%% Load Voltage

t=0:1*10^-10:1*10^-6;

VMax = V0.*Ns;

RTotal = RStage.*Ns;

CTotal = CStage./Ns;

Beta1 = 1./(RTotal.*CLoad);

Beta2 = 1./(RLoad.*CTotal);

VLoad=(VMax/((Beta1-Beta2)*RTotal*CLoad))*(exp(-Beta2*(t))-exp(-Beta1*(t)));

rt = risetime(VLoad)./10^10;

VMaxPeak = max(VLoad)./10^3;

fband = (0.35)./rt;

%Plot VLoad

plot(t.*10^6,VLoad./10^3,'Linewidth',2)

grid on

str = sprintf('Marx Generator Output Voltage for N_{s} = %d and T_{RISE} = %.2fns, V_{peak} = %.2fkV',Ns,rt.*10^9,VMaxPeak);

title(str);

ylabel('V (kV)');

xlabel('Time (\mus)');

%Voltage Sensor

rtvs = rt./5;

fbsens = (0.35)./rtvs;

%% Voltage Divider

%Low voltage arm

Zc = 50;

Z1 = Zc;

Z2 = Zc;

Z = Zc/2;

Tlv = 0;

%High voltage arm

Zhv = (((VMaxPeak*10^3)/80)*Zc)-Zc;

%Resistive Column

ColumnHeight = (VMaxPeak./5).*10^-2;

wire_radius = 12*10^-6;

rmandrel = 0.05;

wire_length = ((pi*(wire_radius)^2)*Zhv)/evan_rho;

wirel = wire_length*2;

Nt = wirel./(pi.*rmandrel.*2);

twirelength = wire_length*4;

%Energy

t=0:1*10^-10:300*10^-6;

VLoad=(VMax/((Beta1-Beta2)*RTotal*CLoad))*(exp(-Beta2*(t))-exp(-Beta1*(t)));

figure

plot(t.*10^6,VLoad./10^3,'Linewidth',2)

grid on

str = sprintf('Marx Generator Output Voltage for N_{s} = %d and T_{RISE} = %.2fns, V_{peak} = %.2fkV',Ns,rt.*10^9,VMaxPeak);

title(str);

ylabel('V (kV)');

xlabel('Time (\mus)');

% x0 = interp1(VLoad,t,0)

%Energy

tmax = 130*10^-6;

Q = ((tmax*(max(VLoad)^2))/(2*Zhv));

%Wire Mass

wire_vol = (pi*(wire_radius^2))*twirelength; %m^3

wire_mass = wire_vol*evan_dens; %kg

%Temperature

deltaT = Q./(wire_mass.*evan_heat)

% Capacitance was 16pF with graded column made of aluminium

cHV = 16*10^-12;

%Connecting Cable

%Chosen wire length of 2m

width = 0.8;

connecting_wire_radius = 5*10^-3;

indSys = indpoly('Rectangle',width,ColumnHeight,connecting_wire_radius)

indConnector = indpoly('RoundWire',width,connecting_wire_radius)/4

%Transfer functiona and response time from notes

rd_sys = 1.2*sqrt(indSys./cHV);

rd_connector = 1.2*sqrt(indConnector./cHV);

Ttotal = sqrt(((rd_sys.*cHV)+(rd_connector.*cHV))^2)

Q2.m

%% Fast Tansient Sensors - Q2 - Coursework 3

% B126949 - Tom Young

%% Pre - Cursor

clear all

clc

syms x

%% Constants

%Global constants

u0 = 4*pi*10^-7;  % H/m

e0 = 8.85*10^-12; % F/m

%Bank constants

Cb = 1*10^-6;

Rb = 10*10^-3;

Lb = 50*10^-9;

Vo = 40*10^3;

%Rod Diameter

Rd = 10*10^-3;

%% Current Discharge

t = 0:1*10^-10:25*10^-6;

[Current_Discharge, Diff_Current_Discharge, damping_string] = dcdisch(Vo,Lb,Cb,Rb,t);

figure('name','RLC Current Discharge')

plot(t.*10^6,Current_Discharge.*10^-3,'Linewidth',2)

grid on

xlabel('Time (\mus)')

ylabel('Current (kA)')

title(['Current Discharge for an ', damping_string ,' RLC Circuit'])

figure('name','RLC Differential Current Discharge')

plot(t.*10^6,Diff_Current_Discharge.*10^-9,'Linewidth',2)

grid on

xlabel('Time (\mus)')

ylabel('Current/Time (GA/s)')

title(['Differential Current Discharge for an ', damping_string ,' RLC Circuit'])

%% Probe Design

%Coil Parameters

major_radius = 20*10^-3

minor_radius = 2*10^-3

wire_radius = 1*10^-3

wire_diameter = 2*wire_radius;

N = 10 %5-10 turns

copper_rho = 1.7*10^-8;

%Current Viewing Resistor

Rcvr = 50 %matches cable impedence

%Pitch

p = (2.*pi.*major_radius)./N;

%Inductance

Induct_pre = u0.*minor_radius.*N;

Induct_sum0 = 0.0007*((log((2*pi*major_radius)/p))^0);

Induct_sum1 = 0.1773*((log((2*pi*major_radius)/p))^1);

Induct_sum2 = -0.0322*((log((2*pi*major_radius)/p))^2);

Induct_sum3 = 0.00197*((log((2*pi*major_radius)/p))^3);

Induct_sum = Induct_sum0 + Induct_sum1 + Induct_sum2 + Induct_sum3;

Induct_pre2 = (((pi*minor_radius)/p)+(log((2*p)/(wire_diameter)))-(5/4)-Induct_sum);

RCoil_inductance = Induct_pre*Induct_pre2

%Resistance

freq = 1/(2*pi*sqrt(Lb*Cb));

RCoil1 = N/(pi*wire_diameter);

RCoil2 = sqrt((copper_rho*pi*freq*u0)*((p^2)+((2*pi*minor_radius)^2)));

RCoil_resistance = RCoil1*RCoil2

%% Magnetic flux

%Magnetic flux time rate-of-change

k = u0.*N.*(major_radius-sqrt((major_radius^2)-(minor_radius^2)));

flux = k.*Current_Discharge;

figure('name','Magnetic Flux')

plot(t.*10^6,flux.*10^6,'Linewidth',2)

grid on

xlabel('Time (\mus)')

ylabel('Magneic Flux (\muT)')

title(['Magnetic flux for an ', damping_string ,' RLC Circuit'])

% Differential Magnetic Flux

dflux = k.*Diff_Current_Discharge;

figure('name','Differntial Magnetic Flux')

plot(t.*10^6,dflux,'Linewidth',2)

grid on

xlabel('Time (\mus)')

ylabel('Differential Magneic Flux (T/s)')

title(['Magnetic flux for an ', damping_string ,' RLC Circuit'])

%% Rogowski Coil Current

Rrt = RCoil_resistance + Rcvr;

Lr = RCoil_inductance;

omega = sqrt((1./(Lb.*Cb))-((Rb./(2.*Lb))^2));

rogowski_current = @(x) exp((Rrt./Lr).*x).*(k.*(Vo./(omega.*Lb).*((omega.*exp(-1.*(Rb./(2.*Lb)).*x).*cos(omega.*x))-((Rb./(2.*Lb)).*exp(-1.*(Rb./(2.*Lb)).*x).*sin(omega.*x)))));

t2 = 0:1*10^-10:10*10^-9;

RIntegral = zeros(0,length(t2));

for n=1:length(t2)

RIntegral(n) = integral(@(x)rogowski_current(x),0,t2(n));

end

Rogowski_Current = (exp(-(Rrt./Lr).*t2)./Lr).*RIntegral;

dRogowski_Current_prefix = (exp(-(Rrt./Lr).*t2)./Lr).*exp((Rrt./Lr).*t2).*(k.*(Vo./(omega.*Lb).*((omega.*exp(-1.*(Rb./(2.*Lb)).*t2).*cos(omega.*t2))-((Rb./(2.*Lb)).*exp(-1.*(Rb./(2.*Lb)).*t2).*sin(omega.*t2)))));

dRogowski_Current = dRogowski_Current_prefix + (RIntegral.*((-Rrt.*exp((-Rrt./Lr).*t2))./(Lr^2)));

figure('name','Rogowski Coil Voltage')

plot(t2.*10^9,Rogowski_Current,'b','Linewidth',2)

grid on

xlabel('Time (ns)')

ylabel('I_{R} (A)')

title('Rogowski Coil Current')

figure('name','Differential Rogowski Coil Voltage')

plot(t2.*10^9,dRogowski_Current.*10^-9,'b','Linewidth',2)

grid on

xlabel('Time (ns)')

ylabel('dI_{R}/dt (GA/s)')

title('Differential Rogowski Coil Current')

figure('name','Rogowski Coil Characteristics')

hold on

plot(t2.*10^9,Rogowski_Current.*Rrt,'b','Linewidth',2)

plot(t2.*10^9,dRogowski_Current.*Lr,'r','Linewidth',2)

grid on

xlabel('Time (ns)')

ylabel('V_{R} (V)')

title('Rogowski Coil Voltage')

legend('R_{R}I_{R}','L_{R}dI_{R}/dt')

hold off

%% Input Voltage

%Differential Magnetic Flux

Vin = k.*Diff_Current_Discharge;

figure('name','Input Voltage')

plot(t.*10^6,Vin,'Linewidth',2)

grid on

xlabel('Time (\mus)')

ylabel('Voltage (V)')

title('Rogowski Coil Output Voltage')

%% Oscilloscope Voltage

% Integrator Design

Ri = 50*10^-3;

Ci = 500*10^-9;

tau = Ri.*Ci;

Vosc = (Rogowski_Current.*k)./tau;

figure('name','Rogowski Coil Voltage')

plot(t2.*10^9,Vosc,'b','Linewidth',2)

Voutp = max(Vosc);

grid on

xlabel('Time (ns)')

ylabel('Voltage (V)')

str = sprintf('Integrator Circuit Output Voltage with peak: %.1fV',Voutp);

title(str)